SLEEP SURFACES: INDICATIONS FOR USE

LOW RISK (Braden Score 15-18)

FOAM MATTRESS

(Found on brown bed frames)

- Low risk ambulatory patient

- Weight limit up to 300 lbs.

- At risk for “bottoming out” if patient’s weight distribution is in the mid or lower body section

MODERATE RISK (Braden Score 13-14)

ZONEAIRE

(BROWN 850, ADVANCE & ADVANTA-STYLE FRAMES)

- Moderate to high-risk patient who does not require low air loss therapy

- Weight limit up to 300 lbs.

- For treatment of Stage I, II and III - VI healing pressure ulcers on one turning surface

- Automatically adjusts 3 zones within the mattress to patient’s weight distribution + 3 heels zones manually activated by caregiver

- Heel Relief Mode for the patient who is at high risk for pressure ulcer formation on heels

VERSACARE

- Weight based pressure redistribution – superior to ZoneAire

- Weight limit up to 500 lbs.

- Therapeutic up to healing, uncomplicated Stage III to IV pressure ulcer on one turning surface

- Zoned mattress automatically turns on when bed plugged into electric source and mattress power button is turned on.

- In-bedded weights

- Length of bed can be adjusted 75” – 86” to fit patient

- 18” from floor to top of mattress

ACCUMAX CONVERTIBLE  MATTRESS

(FOUND ON VERSACARE AND CARROLL BED FRAMES)

(Bed frame has red label, to be used for low bed option)

- Prevention and treatment for Stage I & II and uncomplicated III & IV pressure ulcers

- Weight limit up to 500 lbs.

- Non-powered dynamic pressure redistribution, can be converted to powered therapy by adding alternating pressure pump, found in SPD or on MAP; unit will then offer Alternating Pressure Therapy

- Sloped heel section

- 20 inches from floor to top of mattress on Carroll bed frame

- 21 inches from floor to top of mattress on VersaCare bed frame

MODERATE (Braden Score 13-14) TO HIGH RISK (Braden Score 10-12) WITH MOISTURE ISSUES

VERSACARE  P500

(Bed frame has green label to indicate low air loss.)

- Low air loss therapy

- Weight limit up to 500 lbs.

- Turn assist feature

- Aids in prevention/treatment of pressure ulcers up to Stage III/IV

- Egress from chair position at foot of bed. May also exit bed from side as with standard bed

- Motor driven “Intellidrive: for powered transport

- 19” from floor to top of mattress

ENVISION MATTRESS (RENTAL)

- Low air loss therapy for excessive moisture control, when our owned beds are all in use

- Step-down surface of choice from Clinitron II and immerses patient into bed surface

- Therapeutic weight limit up to 400 lbs.

- Turn assist feature

- Max Inflate feature to decrease shear and drag during patient transfers

TOTAL CARE SpO2RT THERAPY SYSTEM

(For use in Adult ICU’s only)

- To help with prevention and treatment of pulmonary and skin complications related to immobility

- Therapeutic weight limit up to 460 lbs.

- Removes excess heat and moisture from surface for cooler, dryer skin (microclimate control)

- 30 degree head of bed alarm and trending

- Turn assist feature

- Continuous lateral rotation therapy, percussion and vibration

- Full chair egress or may exit bed from side as with standard bed.

- Retractable foot mechanism

- Integrated scale

- Motor driven “Intellidrive: for powered transport

TOTALCARE BARIATRIC – FOAM

(Bed frame has blue label to indicate bariatric.)

- For management of the low risk ambulatory obese patient

- Weight limit up to 460 lbs.

- Egress from chair position at foot of bed or may exit bed from side as with standard bed

- Adjustable length

- Integrated scale

- Does not accommodate overhead frame and trapeze, some beds may have patient helper

- Motor driven “Intellidrive” for powered transport

TOTALCARE BARIATRIC PLUS

(Bed frame has blue label to indicate bariatric.)

- For management of the moderate to high risk obese patient with low air loss therapy

- Weight limit up to 500 lbs.

- Egress from chair position at foot of bed or may exit bed from side as with standard bed.

- Seat deflate option for low bed height when getting in and out of bed

- 40” wide turning surface

- Adjustable length

- Turn assist feature

- Motor driven “Intellidrive” for powered transport

EXCEL BARIATRIC BED

(Bed frame has blue label to indicate bariatric.)

- Bariatric Foam Mattress (Low air loss option for high risk & is a rental item)

- Weight limit up to 1000 lbs.

- Trapeze specific to bed

- Integrated scale

- Not motorized

- Adjustable length

- Seat and foot deflate for low bed height

- Does NOT sit in chair position. Patient egress as with standard bed.

- Bed width can expand from 40” – 50” to accommodate patients with wider girth

VERY HIGH RISK (Braden Score of 9 or lower)

CLINITRON RITE HITE (RENTAL)

- Provides combination low air loss and air fluidized therapies

- Treatment of Stage III or IV pressure ulcers, flaps, grafts, burns and lower leg/trunk wounds, excessive maceration, intractable pain

- Weight limit up to 350 lbs.

- HOB can be raised

- Trapeze can be added

- Easier exit and transfers

- Bed does not come with a scale

- Cannot put patient in trendelenburg and reverse trendelenburg positions

CLINITRON II (RENTAL)

- High air flow

- Treatment of Stage III or IV pressure ulcers, flaps, grafts, full body burns, multiple areas of skin breakdown; ≤ 2 turning surfaces, Intractable pain, severe contractures with breakdown that has occurred on all previously tried surfaces

- Weight limit  up to 215 lbs.

- Eliminate pressure and maceration

- Minimizes friction and shear

- HOB elevated with foam wedges only

- Trapeze available

- Wounds must be covered with occlusive dressing to maintain moisture

NOTE: Patient on Clinitron is at risk for dehydration also pneumonia if poor cough effort.